+24VDC

(From general power supply: Lamp V1 Chain_OK)

+24V_Refill

Pwr IN

Temp

Controller

Trap Nitro. Level

HI contact

Rel31

Rel30.1

Rel31.1

STOP

REFILL

Red Lamp

+24VDC

0V

AL2

Rel30

Rel_24V

Rel_24V.1

Red Lamp

+24V_Refill

220V AC

Rel_24V.2

0V

0V

Rel_Gen

Rel_Gen.1

To Level control instrument

power supply

Power Supply +24VDC _ 6A   DIN format

+24V_Refill

Green Lamp

Refill started… Waiting for LN

0V

VR1

VR2

VR3

VR4

VR5

Rel31.2

Rel_T1.1

+24V_Refill

Rel_LN.2

Rel_LN.1

Temp

Controller

0V

Probe PT100 T1

AL1

Rel_T1.2

Rel_T1

Green Lamp

Refilling Cryotrap

Rel_LN

VR1: 24V valve normally open: security output of dewar when no operation – to close at the beginning of refill

VR2: 24V valve normally close: input compressed air – to open at the beginning of refill

VR3: 24V valve normally close: output LN for cryotrap – to open at the beginning of refill

VR4: 24V valve normally open: output gases at the beginning of refill – to close when refilling if PT100 = -196 deg

VR5: 24V valve normally close: input LN in the cryotrap – to close when refilling if PT100 = -196 deg

- Manual action:  press ‘START REFILL’, and then

  - Ensure that gas outputs at the beginning then valves VR4 and VR5 are activated and LN outputs in cryotrap.

  - The ‘Trap Nitro Level HI’ contact should stop the process; otherwise press ‘STOP’

- Automatic (by remote): send command REFILL (release the command), and then

  - Monitor the T1 PT100 probe temperature, and then monitor the LN level.

  - The ‘Trap Nitro Level HI’ contact should stop the process; otherwise send command ‘STOP’

Ver F (24/08/210) : add relay REL_Gen to reverse the ‘Trap Nitro Level HI contact’ information.

The instrument (AMI – model 186 Liquid Level Instrument) tries to keep the level between the HI and Lo values (=> output energized). By default the internal ‘Hi level relay contacts’ are closed and when the level reaches the HI value they open so it’s necessary to invert the information.