+24V IN

KEY_General Power

To Temperature Controller power supply

0V

Lamp +24V_IN

+24VDC

+24VDC

S

X

Trap Nitro. Level HI contact

Trap Nitro. Level LO contact

V3 Status (relayed)

Lamp V1

Chain_OK

0V

Temp relay (5s)

Rt1 (see p2)

Rlevel: powered if something is wrong

Trap Nitro. Level

0V

Temp. Control. I6

V3 Status

Trap Nitro. Temp (T1 + T2 + T3 + T4)

+24VDC

Temp. Control. I7

Air Pressure

Air pressure (relayed)

0V

0V

Temp. Control. H2

0V

KEY_V2

MANUAL KEY

MaxiGauge P1_2

MaxiGauge P1_1

MaxiGauge P2_2

Temp. Control. H1

MaxiGauge P2_1

MaxiGauge P3_2

Lamp Maxigauge-2

Temp. Control. I8

0V

MaxiGauge P3_1

Lamp Maxigauge-1

0V

KEY_V1

MANUAL KEY

R21.1

R10.1

R20.1

0V

OPEN V1

OPEN V2

(Electrovalve powered)

R11.1

(Electrovalve powered)

CLOSE V2

R20

V1_Powered

V1

V2_Powered

V2

R10.2

R20.2

(Electrovalve released)

CLOSE V1

R10

R11

R21

R11.2

R21.2

(Electrovalve released)

- Schematic Design

+24V IN

KEY_General Power

+24VDC

Lamp +24V_IN

0V

0V

Contacts

Relay Command V1

Contacts

Relay Command V1

(Electrovalve powered)

(Electrovalve released)

(Electrovalve powered)

OPEN V1

Lamp

V1_Powered

0V

V1

OPEN V2

Lamp Maxigauge-1

Maxigauge Contacts:

3 pressures monitored

Contacts

Relay Command V2

Contacts

Relay Command V2

Lamp

V2_Powered

V2

Security Contacts V1Chain_OK:

- liquid Level LO and HI

- Temperatures (4 probes)

- Air Pressure

- V3 valve closed

Lamp V1

Chain_OK

0V

Maxigauge Contacts:

3 pressures monitored

KEY_V1

Lamp Maxigauge-2

KEY_V2

0V

Relay Command V2

Relay Command V1

CLOSE V1

CLOSE V2

(Electrovalve released)

By software the sensors and the outputs of the Maxigauge are linked as following:

| Schematic | Maxigauge | Sensor |
| --- | --- | --- |
| MaxiGauge P1_1 | MaxiGauge relay A | Sensor 1 |
| MaxiGauge P2_1 | MaxiGauge relay B | Sensor 2 |
| MaxiGauge P3_1 | MaxiGauge relay C | Sensor 3 |
| MaxiGauge P1_2 | MaxiGauge relay D | Sensor 1 |
| MaxiGauge P2_2 | MaxiGauge relay E | Sensor 2 |
| MaxiGauge P3_2 | MaxiGauge relay F | Sensor 3 |

Rev G

- Add the Rlevel relay in order to reverse the information from Liquid Level Controller.

. If the power of the Controller is off: HI and LO contacts are closed => Rlevel is powered and the contact is opened.

. If one of the levels of the Controller is out of setpoint: HI or LO contact is closed => Rlevel is powered and the contact is opened.

. If both levels of the Controller are in the setpoint: HI and LO contacts are opened => Rlevel released and the contact is closed.

- Cabling of Rt1 (temporized relay): coil on the +24VDC and contact chain (V3status – Trap Nitro Temp - Trap Nitro Level – Air Pressure) between X and S (Set).

| In order to have a temporization on Rt1 release, following switch configuration has been put : |  |  |  |
| --- | --- | --- | --- |
| █; 6. █; 5. █; 4. █; 3. █; 2. █; 1. |  | Temporization : from 1 to 10 seconds (to be precised on the relay) |  |
|  |  | Mode : delay on release |  |

Rev H

- Add two relays (on V3 status and Air Pressure to allow the use of one contact for V1 Chain_ok and the information +24DC for acquisition by the temperature controller.

Cabling of the HE-XE102 Temperature Controller I/O (see annex A):

| V1 opened | DI1 | J1.1 | J2.1 | C6 |  |
| --- | --- | --- | --- | --- | --- |
| V1 closed | DI2 | J1.2 | J2.2 | R6 |  |
| V2 opened | DI3 | J1.3 | J2.3 | C5 |  |
| V2 closed | DI4 | J1.4 | J2.4 | R5 |  |
| V3 opened | DI5 | J1.5 | J2.5 | C4 | Trap Nitro Temp. Contact 2 |
| V3 closed | DI6 | J1.6 | J2.6 | R4 | J2.7 Link |
| Air Pressure OK | DI7 | J1.7 | J2.7 | C3 | J2.6 |
| MaxiGauge1 OK | DI8 | J1.8 | J2.8 | R3 | J2.9 Link |
| MaxiGauge2 OK | H1 | J1.9 | J2.9 | C2 | J2.8 |
| GND | 0V | J1.10 | J2.10 | R2 | J2.11 Link |
| Temp Probe #1 | AI1 | J1.11 | J2.11 | C1 | J2.10 |
| Temp Probe #2 | AI2 | J1.12 | J2.12 | R1 | Trap Nitro Temp. Contact 1 |
| Temp Probe #3 | AI3 | J1.13 | J2.13 | H4 |  |
| Temp Probe #4 | AI4 | J1.14 | J2.14 | H3 |  |
| Temp Probe Com | Com | J1.15 | J2.15 | H2 | V1 Chain_OK |

Annex A: Temperature Controller

| V1 opened V1 closed V2 opened V2 closed V3 opened V3 closed Air pressure MaxiGauge1 Maxigauge2 GND Temp probe 1 Temp probe 2 Temp probe 3 Temp probe 4 Temp probe Com |  | V1 Chain_OK Contact probe 1 Contact probe 2 Contact probe 3 Contact probe 4 4 temperature probes from the trap : |
| --- | --- | --- |

Annex B: Cryo Trap

Valve V1

Valve V2

Valve V3

Air for open/close the valve

Valve status cable