5.   Scattered light noise

5.1 Introduction

The sensitivity of the interferometer can be limited by the noise due to scattered light propagation and recombination in the interferometer beams.

The very high quality mirrors, with a very small coating diffusion scatter some ppm of the stored power, The resulting scattered light propagates in the vacuum chamber, and after reflections and diffusion from the walls, may recombine with the main beam, with a different phase due to the different optical path. The vibrations of the chamber makes the phase shift time dependent, resulting in noise in the phase of the interference signal.

5.2 Noise evaluation

The noise was evaluated, in order to appreciate the necessity of reducing it, then to compare the efficiency of possible ``killing’’ systems, with baffles. The initial ideas of this evaluation have been formulated by K.Thorne (5.1). We have developed a statistical approach of the scattered light propagation (5.2). The information on trajectories obtained from a particle analysis is supplemented by additional information regarding phase modulation coming from the wave picture. Scattered light was represented by optical rays, extracting information on the noise from the statistics of photons received by the mirrors. Then the options for a baffle system have been evaluated by a Monte Carlo simulation. Coherent effects, such as diffraction or back reflection by the baffle edges, were computed analytically (5.3).

As a first result, the noise in an empty tube has been evaluated to be h(f) ~10-23/Hz. Being of the same order as the foreseen sensitivity, it is obvious that a system of baffles had to be found and implemented in the arm tubes. Calculations of the same kind have also shown that scattered light has to be attenuated in the central area.

5.3 Baffles in the arm tubes

It was decided to allow the possibility of having three interferometers side by side inside the vacuum tank. For this a free aperture of 1 m is necessary, which leaves a space in the tube cylinder between  1000 mm and  1200 mm available for baffles.

The best performing material was absorbing black glass for welding protection  with an anti-reflective coating (Schott-Desag Athermal 11, thickness 3.2 mm, transmission ~ 10-7, reflection between 3.10-3 and 10-2 depending on incidence angle, negligible diffusion, at  = 1064 nm). A satisfactory glass prototype was realized, with the shape of a quasi perfect light trap, but finally these baffles were found too fragile for installing in the 3 km tubes. The black glass was used only in the central area and in the links between valves and towers. The baffles in the 3 km arm tubes have been built in stainless steel.

5.3.1 Steel baffles

Clearly steel baffles with a conic shape are more efficient than simple rings normal to the tube axis. Stainless steel has a small specular reflection coefficient, but there is a large diffusion angle distribution function (BRDF), and the absorption is close to 50%. The BRDF has been measured for the sandblasted stainless steel used for the baffles, and was introduced in the Monte Carlo computation.

Finally, steel baffles have the shape of a truncated cone with an external diameter of 1184 mm, internal diameter of 990 mm and an angle of 55° with respect to the cone axis. The steel thickness is 1.5 mm. The internal edge is serrated in order to suppress diffraction by edges. Fig.5.1 shows the cone shape, and the way they are installed in the tubes. The conceptual requirement was to cover completely the solid angle as seen by the beam spot at each Fabry-Perot mirror. A simple calculation gives the distance Zn of the nth baffle to the mirror through the law:

Zn = Z1 αrr−hn−1

Where r is the tube radius, h is the baffle height, and  is a coefficient for getting some coverage between the shadows of each baffle. The total number of steel baffles is 160 per arm. They have been realized by several firms near Paris and soldered inside the modules at CNIM before the washing, by mean of 4 small lugs on their external part. They have been installed only in the tube part in between the two large valves, which is supposed to remain untouched hopefully for the Virgo lifetime.

5.3.2 Glass baffle

In the links joining the large valves and the towers, large glass baffles have been installed for covering all discontinuities like valve or tower flanges, bellows etc... In this area glass baffles are allowed, since they can be relatively easily replaced, in case of breaking. All these link baffles are 1160 mm diameter disks made out of black glass, with a large central hole. Inside each link, there are 4 rings, one in front of each flange (tower and valve), and one on each side of the bellows. The inner hole diameter ranges from 600 mm to 850 mm according to the baffle location  and  the laser beam size. Their sizes and positions are displayed in fig 5.2

With these steel and glass baffles, the evaluated noise in the arm tubes is 
h = 6.5 10-25/Hz.

5.4 Baffles in the Central Area

The central area is constituted by 5 towers, their mirrors (fig 5.2) and the 4 junction links (small links). To evaluate the noise with a Monte Carlo simulation, all possible processes for rescattering of photons off all elements have been listed. The geometry of the tanks, links and mirrors, and the optical properties of all surfaces, have been taken into account (5.4). Photons coming from the main beam have been generated and their scattering off each element has been followed in all their possible paths until a possible recombination. The total noise is the incoherent sum of the contributions of all scattered light generation processes through all recombination processes.

The resulting actions for reducing the noise are given below together with the final result.

- For the small links, the steel tube diameter must not be less than 400 mm, to avoid the effects of grazing diffusion on its walls. Black glass disks with external diameter  380 mm, and a central hole of diameter 230 mm, corresponding to ± 5 beam waists, have been inserted inside the links. Their position was chosen for hiding the 250 mm diameter valves and the bellows waves. With these baffles (14 altogether) the noise coming from the small links is negligible.

- Among the towers the most delicate situations can be found around the two input mirrors and the beam splitter, due to the complex geometry and the high light intensity.  In fact all possible diffusion or reflection areas in the tanks have to be covered with black glass, with ± 5 waist holes for beam passage. In order to be as close as possible to the tower cylinder, the baffles inside the towers at the junction of the big links have the shape shown in Fig.5.3: in the centre a big plate 1150  600 mm2 with a ± 5 waist hole, and on each side three 1150  120 mm2 plates installed as shown in the side view in Fig.5.3. The beam splitter tower has 4 such devices, covering most of the cylinder (Fig.5.4), the input towers have 2, the end towers have one and the recycling tower has one also.

- The resulting noise from the central area is reduced to ~ some 10-25/Hz.

References

5.1. Light scattering and proposed baffle configuration, KS.Thorne  Caltech Report,1989

5.2 Scattered light noise in gravitational interferometric detectors: A statistical approach ,

J-Y.Vinet, V.Brisson, S.Braccini, I.Ferrante, L.Pinard, F.Bondu, and E.Tournié

Phys.Rev. D 56  6085 ,1997

5.3 Scattered light noise in gravitational wave interferometric detectors : Coherent effects

J-Y.Vinet ,V.Brisson ,S.Braccini ,Phys Rev D  54 1276 ,1996

5.4 Diffused light in the central area, S.Braccini, JY.Vinet, and V.Brisson  Virgo note 1995

Figures

5.1     a) Shape of a conic metal baffle with serrations at the inner diameter

b) Implementation of steel baffles inside the 3 km tube

5.2     Scheme of the central area and location of glass baffles

5.3     Shape of a glass baffle covering the tower inner wall around the laser beam path

5.4     Picture of the 4 glass baffles inside the beam splitter tower taken from below with a

fish-eye lens before mirror installation.