Minutes of the VAC Meeting, Tue 21Apr09, 14h30 - 17h

---------------------------------------------------------------------

Presents: C.Bradaschia, F.Carbognani, F.Cavalier, R.Cavalieri, R.Cosci, G.Gemme, A.Paoli, A.Pasqualetti, P.Popolizio,

1 Introduction

1.1 updated work breakdown wbs.xls (with cost quality) available on workarea, to be checked.

1.2 Optical design 'baseline' options: we have to exploit the vacuum system changes and the towers

accessibility in those scenarios. Maybe 3 or 5 towers will be joined together in one unique vauum chamber.

The procedure of opening towers would be much affected. Ot is to be checked wether 800mm valves can be installed

in between some towers.

1.3 Pumping system of central area: ‘scroll‘ will be always ON but displaced in a suitable room, to cope with env.noise.

1.4 Ion pumps will be employed to pumps the mirror chambers (IB, DT, MC. No turbos on towers lower parts).

1.5 Perhaps small traps are feasible on bench towers links:

depending on optical design the diameters may be similar to the present one.

1.6 Baseline design: document of about 7-8 pages to be prepared within the 8May. Aim is to describe the status of vac at the end of the external review process.

2 LN2: Criotech company quoted as 3%-5% the daily LN2 consumption of the standard dewars for rent.

Considering the traps consumption , about 150liter/day, One 10000liters tank may shall be not

sufficient to guarantee > 2weeks of uninterrupted operation.This seems not acceptable.

Furthermore the cost for LN2 would be considerable. This situation is not a probelm for the normal

users of LN2 because they make use of the gas N2, evaporating the LN2.

Informations on supply contracts from different companies will be taken to assess the situation.

3 Cryotraps: not discussed. a 'cryotrap workshop' shall be prepared in the next weeks (may or early june),

to discuss design details

4 UHV air inside towers: not discussed. Pay design is starting, we shall follow actively its advancement

5 Control system:

Architecture basics: 1 control racks per pumping station (= 1 tower or 1 tube pumping station, totally about 30 units), with possibility to operate locally.

For the control racks , there are two general options for the logic: make use of PLC or of industrial PC, with different features. In particular the PLC option is the preferred one concerning the , i.e. less SW, but is not well adapted to manage a large number of RS232 devices (about 10 per tower). Alternative to RS232 be indagated: for instance RS485, profibus should be ok for PLCs.

RS232 has also the drawback of a limited cable length, possibly a problem for racks displacement

Maybe the ethernet /serial converters could be a solution to interface PLCs with serial devices, to be checked, in particular for reliability .

The cost has been revised, it is to be upgraded at 450keur

6 HW & SW integration

Advantages of general integration of he HW & SW of the different systems: vacuum, suspension controls, optical devices,  infrastructure, …also concerning possible common solution for EM reduction like racks displacement, mains supply,  . Strict interface with DAQ subsystem . Dedicated discussions, involving all concerned people, will take place in the following months.

7 Tower displacement:

The displacement of selecteed central area towers involves to dismount the platform parts . Once towers wil be displaced in the new positions,  new platform pasrts, with a different size, shall be re-installed.

It is valuable to have a less ‘noisy’ platform for AdV: presently, the people presence on the platform durinq commissioning works produces the unlock of the itf. It is worth to correct this trouble.

Also the present ‘towers scaffoldings’ have to be changed with new ones for given reasons (P.Popolizio, F.Richard).

A new platform-design , without large vibrating parts exposed to the air conditioning flux, would be also positive concerning the environmental noise in the central hall.

We are currently exploiting the displacement work details; in some time we shall call for discussions among platform users and collaboration members to set up ‘new platform’ design.

8 vac compatibility test.

LAL may help ego to perform vacuum tests of materials, (not more than 2-3 per year) to be confirmed.

This activity could be shared between ego and lal . Specific tasks are mentioned in the various

subsystems (INJ,PAY,DET,SAT,TCS), presently not repeated in the vac wbs.

9 diffused light

still open point = cryotrap minimum aperture . Valve size to be defined (650mm or 800mm). It needs an evaluation of the diffused light effects on the itf in both design options.

Experts (C.Bradaschia, V.Brisson, JYVinet ) should provide answers within end of 2009.

Actions summary:

----------------------

1) define the ideas about the 'new platform', and prepare a presentation. At first at a AdV biweekly, then at the collaboration . Roberto, Frederic, Tatiana

2) provide a list of the Rs232 devices in a tower,  Lal team

3) investigate the possibility of changing the RS232 interfaces on existing equipment, EGO vacuum group

4) For each task: Verify and check the wbs and planning documents on the web

5) All: Contribute to the ‘baseline design’ document. A first version is in preparation starting from the existing documents (‘preliminary design’and the various ‘reviews documents’ ) ; it will be circulated in the next week for comments and corrections. Normally half a page for each task will be needed.

6) cryogenic experts,+Rapagnani, Ricci,...: Prepare a design meeting on Cryotrap within early June.