List of electrically connected devices inside a long tower

SATCS meeting   -  14feb2012   -    V.Dattilo

|  | TopStage TopRing BottomRing | F#0 | F#1 F#2 F#3 F#4 | F#7 | Marionette | Mirror C.Plate | TCS | Tower Bottom | TOTAL | TOTAL Number of pins |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| MOTORs | 6 | 4 | 4 | 5 | 1 (2) |  | 3 |  | 23 (24) | 161+46 (168+48) |
| COILs | 3 | 2 |  | 3 (6) | 4 (8) | 4 |  |  | 16 (22) | 32+16 (44+22) |
| ACCELEROMETERs | 3 | 2 |  |  |  |  |  |  | 5 | 30+15 |
| LVDTs | 3 (6) | 1 | 4 | 4 (7) |  |  |  |  | 12 (18) | 48+24 (72+36) |
| TEMPERATURE sensors |  | 2 | 8 | 2 | (1) | (1) |  | 2 | 14 (16) | 28+14 (32+16) |
| PIEZOS | (3) |  |  |  |  | (3) |  |  | (6) | (12+6) |
| TILTMETERS | (3?) |  |  | 2 (0) |  |  |  |  | 2 (3?) | 6 (?) |
| CHRoCC heater |  |  |  |  |  |  |  | 1 | 1 | 2 |
| CHRoCC T sensors |  |  |  |  |  |  |  | 4 | 4 | 8 |
| HEATING RING |  |  |  |  |  | (2) |  |  | (2) | (4+2) |
| TOTAL (~upper limits) |  |  |  |  |  |  |  |  | 77 ( 101) | 315+115 (378+145) |

-In red is indicated the number of devices (or pins) for AdV, only in case it is changed with respect to Virgo+

- the total number of feedthrough pins is indicated as  a+b,   where a=pins of the device,  b=pins for shields

ELECTRICAL FEEDTHROUGHS

- In Virgo+ Feedthroughs for long towers are available in:

2 x CF250 flanges (11 feedthroughs x 32 pins - MIL-C-26482), mounted on technical virole, a few feedthroughs are not used (available as spares)

1 x CF160 flange (2 feedthroughs x 32 pins - MIL-C-26482), mounted on tower lower part  (for CHRoCC and TCS)

- For AdV:

Keep the 2 x CF250 flanges;

Add feedthroughs on tower lower part (for F7 LVDT/actuators)

One vacuum-side cable connected to each feedtrough.

A single cable connects devices placed on the same filter or suspensions place  a single cable connected to different devices (for instance 1 motor, 1 LVDT, 2 temperature sensors).

Also for AdV it is not foreseen to use the single cable to connect devices of the same kind.

Slight differences of feedthrough pinout among long towers will be still possible.